## Slide 1: Introduction to Neutron Radiography and Computed Tomography
Neutron Imaging Team at CG-1D:
Indu Dhiman, Postdoc Research Associate
Lou Santodonato, Scientific Associate Jean Bilheux, User Imaging Software Data Analysis Jiao Lin, Reconstruction Steven Hahn, 3D visualization Gian Song, Postdoc Research Associate Barton Bailey, Engineer Hassina Bilheux, Instrument Scientist
NX School August 5, 2016

## Slide 2: Neutrons Measure Structure
10-11
10-9
10-7
10-5
10-3
Dimension (meters)
0.1Å
1.0nm
1mm
0.1mm
10.0mm
Neutron Diffraction
Neutron Scattering
Neutron Microscopy
Neutron Imaging
Inferred structure (indirect)
Direct structure
Nuclear and electronic density in enzymes
SANS used to construct protein kinase A (PKA)
Fluid interactions in plant-groundwater systems
Soil-root interface (rhizosphere)
In Vivo Study of Embolism Formation
Computed tomography
Characterization of biological membranes, colloids, porosity, etc.
Neutron diffraction of D2 sorption in Cu3[Co(CN)6]2
Ice/water segregation in permafrost structures

## Slide 3: Neutron sensitivity
[M. Strobl et al., J. Phys. D: Appl. Phys. 42 (2009) 243001]
Neutron Radiograph of camera
X-ray Radiograph of camera
Courtesy of E. Lehmann,PSI
Neutron Radiograph of Rose in Lead Flask

## Slide 4: Imaging throughout Nobel Prize history
1901: Roentgen, FIRST Nobel Prize in Physics, Discovery of X-rays
1979: Cormack and Hounsfield, Nobel Prize in Medicine, Computed Tomography (CT)
1986: Ruska, Binnig, Rohrer, Nobel Prize in Physics, Electron Miscroscopy
2003: Lauterbur and Mansfield, Nobel Prize in Medicine, Magnetic Resonance Imaging (MRI)
2009: Boyle and Smith, Nobel Prize in Physics, Imaging semi-conductor circuit, the CCD* sensor
(*) Charge-Coupled Device

## Slide 5: What is imaging?
Imaging is the visual representation of an object: photography, cinematography, medical imaging, X-ray imaging, thermal imaging, molecular imaging, neutron imaging, etc.
Digital Imaging is a field of computer science covering images that can be stored on a computer as bit-mapped images

## Slide 6: Early neutron imaging measurements
Neutron Imaging started in the mid 1930’s but only during the past 30 years has it come to the forefront of non-destructive testing
World class dedicated imaging user facilities such as NIST, PSI, HZB, FRM-II and at many worldwide universities
World conferences and workshops being held regularly
Growing worldwide user community
Discovery of neutron in 1932 by Chadwick
First neutron radiograph in 1935
Left to right: Pressure gauge with metal backplate; fire hydrant and test tubes filled with H2O
and D2O imaged with gamma-rays (top) and neutrons (bottom)
[ Kallman and Kuhn, Research 1, 254 (1947) ]

## Slide 7: Multiple scattering and low detector spatial resolution
[ J. Anderson et al., Br. J. Radiol. 37, 957 (1964) ]

## Slide 8: Today: Comparison microscopy/microCT and neutron radiography
92% of the pixel intensities agreement between histological and neutron
Watkin, K. et al., Neutron Imaging and Applications, Springer, 2009.

## Slide 9: Attenuation-based neutron imaging
Beam attenuation caused by a homogeneous uniformly thick sample composed of a single isotope is given by
t() = scattering and absorption
 is the attenuation coefficient and x is the thickness of the sample
t() is the material’s total cross section for neutrons,  is its density, NA is Avogadro’s number, and M is the molar mass.
Neutron
Radiograph
Photograph

## Slide 10: Detection of “imaging” neutrons
Scintillator-based techniques such as 6Li(n,α) 3H
  Good signal-to-noise (SNR) ratio
  Large Field Of View (FOV) and 0.01 to hundreds of seconds images
  BUT spatial resolution limited by the dissipation of particles
  Can take a lot of neutron flux!
1,1 | 1,2 | 1,3 | … | … | 1, ny
2,1 | 2,2 | 2,3 | … | … | 2, ny
3,1 | 3,2 | … | … | … | 3, ny
… | … | … | … | … | …, nx
nx, 1 | nx, 2 | nx, 3 | … | … | nx, ny
Each pixel is coded using n-bit.
16-bit = pixel value is between 0 and 65535
Notes: n + 6Li  4He + 3H + 4.79 MeV
Collisions with these charged particles cause the ZnS(Ag) to scintillate at a wavelength of approximately 450 nm

## Slide 11: Detection of “imaging” neutrons (cont’d)
Pixelated detectors
  Micro-Channel Plate (MCP)
  In the direct path of the beam
  Limited FOV for high spatial resolution MCPs
    2.8 cm x 2.8 cm at ~ 25-30 microns (in centroiding mode)
  Encodes events at x, y position and time of arrival, at high temporal resolution ~ 1 MHz
  Detection efficiency has improved for both cold (~70%) and thermal (~50%) energy range
  Absence of readout noise
  Not as gamma sensitive (less than 1%)
  Becoming commercial
  BUT: works in relatively low-signal beam!
Notes: n + 155Gd  Gd*  -ray spectrum  conversion electron spectrum
n + 157Gd  Gd*  -ray spectrum  conversion electron spectrum

## Slide 12: Example: Cu, Zn and Al
Chart: Neutron Transmission through Metals
Compound | Abs. Coeff. [Å-2] | Inc. Coeff. [Å-1]
Cu (100%) | 1.78E-09 | 4.65E-10
Al (100%) | 7.75E-11 | 4.94E-12
Zn (100%) | 4.06E-10 | 5.06E-11

## Slide 13: CG-1D Neutron Imaging Facility
Lens
CCD
Scintillator
Mirror
ANDOR Camera:
4Mpixels – 2048x2048
Field of view:
7x7cm2
Quantum Efficiency:
95%

## Slide 14: HFIR Neutron Imaging Facility (CG-1D)
Detector capability/spatial resolution:
ANDOR® CCD Detector ~ 100 microns
sCMOS ~ 50 microns, 100 fps
MCP(*) ~ 100 microns (and ~ 25-30 microns in centroiding mode), or microsecond time frames
Neutron Beam Path
(*) In Collaboration with Prof. Anton Tremsin, UC-Berkeley
Notes: Put a scale on photo
Beam path instead of flight tube
Changed to new photos.

## Slide 15: The CG-1D detector suite and respective common applications
Type of detector | Largest field-of-view (cm2) | Effective pixel size where the image is formed (microns) | Corresponding spatial resolution (microns) | Maximum Dynamic resolution (no pixel binning) | Dynamic range (bit) | Applications
CCD | 7.4 x 7.4 | 36 | ~100 | 1 image/second | 16 | Additive Manufacturing (AM), Energy, Aerospace, Transportation Research, Geosciences, Biology, Botanical, Archeology, Medical, Food Science
sCMOS | 3.3 x 2.8 | 13 | < 100 | 100 images/second for 4 sec max (**) | 16 | AM, Geosciences, Petroleum, Food Science
MCP (*) | 2.8 x 2.8 | 55 | ~100 | 1 million images/second | 14 | Transportation Research, AM, Energy, Aerospace, Biology, Geosciences, Petroleum, Botanical, Food Science
 | 2.8 x 2.8 | 12 | ~ 25-30 | 1 image/ 4 hours | 14 | Biology, Botanical, Medical
(*) The MCP detector has two modes of operation: the high time resolution and the high spatial resolution options, respectively.
(**) When binning pixels or working at lower frame rate, the limit is the 5 TB storage on the sCMOS server (each image is 10 MB). We can thus collect thousands of images.

## Slide 16: CG-1D polychromatic beam
Chart: Wavelength (Å)
CG-1D spectrum measured with the MCP detector at a flight path distance of approximately 5.5 m, with the chopper running at a frequency 40 Hz and an 5 mm aperture. Kang M. et al., Nuclear Instruments and Methods, 708, 2013]

## Slide 17: Use of Diffusers
1 cm thick Graphite Powder (4 to 10 microns)
Open Beam Heat Pipe Sample Reduced Image
Without diffuser
With diffuser
Notes: Now, we use 10 nm Al2O3 nanoparticles enclosed in Al containers (1 mm to 4 mm). Usually, we use the 2 mm diffuser.

## Slide 18: Sample stage for nCT
EGR
Cooler
Neutrons
CCD Detector
Rotation/
Translation
Stage

## Slide 19: Data Normalization for Imaging
2D – Radiography
  Normalization
=
-
-
Image
Dark Field
Dark Field
Open Beam
Normalized Image
1
0
Transmission

## Slide 20: Computed/Computerized Tomography (FBP)
Raw Data:
2048x2048 pixels,
721 projections
Normalized Data:
2048x2048 pixels,
721 projections
Sinograms:
2048x721 pixels,
2048 files
Slices:
2048x2048 pixels,
2048 slices
3D reconstruction:
2048x2048x2048 voxels
  Filtered back projection method
~20700
~700
Counts
1
0
Transmission
1
0
Transmission
∞
0
Attenuation

## Slide 21: Computed/Computerized Tomography (CT)
Several techniques:
  Filtered Back Projection
    Radon transform
    Works well with high signal to noise ration measurements
    Easy-to-use commercial, semi-automated software available
    Quick
  Iterative Reconstruction
    Direct approach
    Less artifacts
    Can reconstruct incomplete data
    High computation time

## Slide 22: Conventional Neutron Imaging Techniques at steady-state sources
Radiography (available at CG-1D)
Tomography (available at CG-1D)
Phase Contrast Imaging
Polarized Neutron Imaging
Stroboscopic Imaging
Imaging of processes that happen fast
Energy selective techniques possible with double-monochromator configuration

## Slide 23: Neutron Imaging Techniques at pulsed sources
Energy-selective (or Time-of-Flight) imaging
  Contrast enhancement
  Bragg edge
Energy resonance imaging
Stroboscopic imaging
  SNS has a natural clock
Neutron Imaging at energies not accessible at reactor facilities
  Mainly bio-medical applications

## Slide 24: Neutron imaging techniques
Radiography
Computed tomography
Bragg edge imaging
Neutron phase imaging
Stroboscopic imaging
Neutron Stimulated Computed Emission Tomography or NSECT
Polarized imaging
Dark field imaging
Energy resonance imaging

## Slide 25: Bragg Edge Imaging
Crystalline structure mapping
Bragg reflected neutrons result in narrow dips in the actual transmission at precise wavelengths specified by Bragg’s law:
hkl = 2dhkl sinhkl
where dhkl is the interplanar distance for the (hkl) planes and hkl are the Bragg angles hkl depends on the relative orientation of the crystal lattice to the neutron beam.
At hkl, creation of map of the number of crystals having any of their (hkl) directions making an angle, hkl, with the incident beam given by:
hkl = (/2) – arcsin (hkl / 2dhkl)
Kockelmann et al., NIM A, Vol. 578 (2007) 421.

## Slide 26: Propagation-based Neutron Phase Imaging
(a) Neutron attenuation radiograph (e) and phase contrast radiograph of a lead sinker mounted on an Al screw. [B. Schillinger et al., Mat. Trans. Proc. (2006) 61]
  Source needs to be spatially coherent (i.e. small pinhole and long pinhole-detector distances)
  Flux is low (up to 98% of flux is sacrificed, several hours to days for one radiograph)
Neutron attenuation radiograph (b) photograph and (c) phase contrast radiograph of a yellow jacket wasp.
[B. E. Allman et al., Nature 408 (2000) 158]
SOURCE
L>>D
D < 1 mm
L > 15 m
Notes: Mention values D, L

## Slide 27: Phase Radiography using Grating Interferometry
G0 creates array of coherent sources from source w
G1 creates diffraction patterns for each source which overlap if
Diffraction pattern has maximum contrast when d is a integer multiple of the Talbot length, LT
Phase object cause distortion of diffraction pattern (or phase shift of incident wave Φ)
Measure diffraction pattern by translating G2
G0: (source) absorption grating, period p0
G1: phase grating, period p1
G2: (analyzer) absorption grating, period p2
[Pfeiffer et al., PRL. 96 (2006) 215505]
Φ
p0 ~ 1 mm, p1 ~ 10 m, p2 ~ 5  m, ~ 5 m, d ~ 20 mm
Notes: ----- Meeting Notes (6/17/11 11:26) -----
This detector has rather poor pixels
G2 allows you to measure this diffraction pattern which has micron sinus waves with a poor resolution detector.
G2 is just a trick for being able to measure intensity with lower resolution detector.

## Slide 28: Stroboscopic imaging
Makes a cyclically moving object appear to be slow moving
Pulsed sources are by definition stroboscopic neutron sources
BMW engine, NEUTROGRAPH, ILL, France
Schillinger et al., NIM A 542 (2005) 142.
Oil spreading into bottom of piston
Oil jet ejected into bottom of piston
Stroboscopic imaging: 150 exposures, 200 ms each, 24 cm x 24 cm field of view

## Slide 29: Polarized neutron imaging
Spin rotation in magnetic field B, length s, v = velocity of neutrons
Setup:
http://iopscience.iop.org/1367-2630/11/4/043013/fulltext/
Magnetically aligned polarized neutron beam is a great probe for the investigation of magnetic materials

## Slide 30: Polarized neutron imaging
Produce a polarized neutron beam.
Preserve the polarization of the beam.
Manipulate the polarization of the neutron beam: if needed.
Deliver it to the sample.
Preserve the change in polarization (if at all), as the beam is transmitted through the sample.
Detect this change using CCD cameras.

## Slide 31
Polarized Neutron imaging set – up on CG-1D

## Slide 32: Applications at a glance
Advanced Manufacturing
Archeology
Bio-medical
Botany
Contraband
Cultural Heritage
Energy
Engineering/Materials Science
Forensic Science
Geology/Earth Sciences
Homeland Security
Paleontology
Quality Assurance
Visualization of water transport in artificial soil sedimentation (20 s frame, 25 x 25 cm2)
http://neutra.web.psi.ch/gallery/animations.html
Radiography of a dry monkey skull
http://neutra.web.psi.ch/gallery/biological.html

## Slide 33: Materials Research, Energy and Engineering Applications

## Slide 34: Neutron Computed Tomography (CT) maps distortion and defects in AM parts
Variance measurements based on comparison between the file used to print and neutron CT
Distortion and defects clearly apparent in the final product
These details can be fed back to modeling tools at Honeywell to optimize printing parameters
=
Notes: Visualization and analysis shown here have been produced with VGStudio Max (Commercial Software).

## Slide 36: Neutron Computed Tomography Characterizes Particulate Properties During Regeneration
Scientific Achievement
  Employed neutron tomography to measure the properties of soot cake layer thickness and subsequently density during a sequential regeneration. Identified key parameters that will aid understanding of the fuel-consuming regeneration process.
  Work performed at the High Flux Isotope Reactor Imaging (CG1D) beamline was supported by Scientific User Facilities Division, Office of Basic Energy Sciences, US Department of Energy.
Significance and Impact
  Understanding the regeneration progression of DPFs is important in improving the efficiency of the process and to understand the correlation between pressure drop and soot level. The results from this work can directly provide model parameters to industry so they can better predict the level soot/particulate in the DPFs.
Research Details
  DPFs were filled to different levels (3, 5 and 7 g/L) using diesel engine
  Full DPFs analyzed with neutron tomography to determine the soot cake thickness and packing density
  Soot cake density, thickness and axial profile measured during sequential regeneration
  Highest soot cake density observed during initial 20% regeneration; layer then loses density, porosity increases and channels open up
  Different soot loading displayed same behavior during regeneration
T. J. Toops, H. Bilheux, S. Voisin, J. Gregor, L. Walker, A. Strzelec, C. E. A. Finney, J. A. Pihl, Nuclear Instruments and Methods in Physics Research Section A 729 (2013) 581-588.
T. J. Toops, J. A. Pihl, C. E. A. Finney, J. Gregor, H. Bilheux, Emission Control Science and Technology 1:1 (2015) 24-31.
DPFs collect soot during engine operation, leading to soot cake. During regeneration soot cake properties change significantly as identified in neutron imaging campaign.
  Diesel particulate filters (DPF) clean exhaust and reduce pollution

## Slide 37: Non-destructive neutron imaging being employed to investigate automotive devices and functionality
Cross-sectional photograph (left) and equivalent neutron radiograph (right)
Unique and powerful neutron capabilities available at ORNL
Neutron imaging is a non-destructive technique and analytical tool
  Particulate distributions
  Quantification possible
  Average distribution and depth for 3 particulate loadings (3, 5, 7 g/L)
Intra-fuel injector flow patterns
  More sensitive to fuel the metal
Sophisticated image processing allows device tomography
1 inch
T. Toops (PI), S. Voisin, C. Finney, J. Gregor, H. Bilheux

## Slide 38: Comparison of Li Distribution as a function of depth of battery
Control
Li-Air (no catalyst)
Li-Air (with catalyst)
Nanda et al., Journal of Physical Chemistry C, 2012.
Notes: All authors listed and title of the publications, anytime paper cited

## Slide 39
Neutron Imaging Provides the Basis for Developing Models
Neutron image
Li- air cathode
3D model
Li-air cathode
Reaction phase 3 dimensional modeling was used to predict results and compare with measurements
Spatiotemporal reaction phase three-dimensional modeling of the electrodes also predicted a non-uniform lithium product distribution, confirming the neutron imaging result.
Need to match resolution of neutron imaging capabilities to further improve feedback to modeling tools
Non-uniform lithium distribution may limit rechargeability

## Slide 40: Rapid Imbibition of Water in Fractures within Unsaturated Sedimentary Rock
Scientific Achievement
  Spontaneous imbibition of liquids into gas-filled fractures in variably-saturated porous media is important in a variety of engineering and geological contexts. Dynamic neutron radiography was applied to directly quantify this phenomenon in terms of sorptivity and dispersion coefficients.
  Work performed at the High Flux Isotope Reactor Imaging beam line (CG1D) was supported by the Scientific User Facilities Division, Office of Basic Energy Sciences, U.S. Department of Energy.
Significance and Impact
  The theory derived describes rapid early-time water movement into air-filled fractures in sedimentary rock. Both theory and observations indicate that fractures significantly increase spontaneous imbibition and dispersion of the wetting front in unsaturated sedimentary rocks. Capillary action appears to be supplemented by surface spreading on rough fracture faces. The findings can be applied in modeling hydraulic fracturing.
Research Details
  Imbibition into unsaturated Berea sandstone cores was controlled with a Mariotte bottle setup. Images were collected using dynamic neutron radiography and analyzed using ImageJ.
Cheng C. -L., Perfect E., Donnelly B., Bilheux H. Z., Tremsin A. S., McKay L. D., DiStefano V. H., Cai J. C., Santodonato L. J., Rapid imbibition of water in fractures within unsaturated sedimentary rock. 2015. Advances in Water Resources, Volume 77, Pages 82–89 <http://dx.doi.org/10.1016/j.advwatres.2015.01.010>
Time sequence of neutron radiographs showing the rapid uptake of water into a longitudinal, air-filled fracture zone in Berea sandstone (~50 mD). FOV is ~28 x 28 mm2.

## Slide 41
Significance
Water in porous media
  agricultural crop production
  infiltration & aquifer recharge
  remediation of contaminated soils
  subsurface carbon sequestration
  enhanced oil recovery
Contaminant transport
CO2 sequestration
Plant-soil-water system

## Slide 42: Neutron Radiography of Roots at CG1-D
Water injected into root zone at base
Unidentified endophyte (symbiotic) or decomposer fungi visible near roots of switchgrass (left), revealing substantial hydration of the rhizosphere
Both fine and coarse roots are readily visible
Switchgrass and Maize
Plants in Al chamber
Notes: Fairly dry condition
Endophytics means roots of fungus spread and attached to roots, increase surface area of roots, able to maintain hydration of roots in the rhizosphere
Bulk water content measured, no idea what happens in the roots and around the rhizosphere

## Slide 43: Changes in Soil and Root Water Content using Neutron Radiography
Notes: Corn, dynamics measurements over 12 hours
Larger root is more hydrated (white color), dark areas means water extracted from system (less hydrated roots)
Loosing conductivity around the smaller roots
Water content increases in the top of circle over time: root growth and water release

## Slide 44: Water Uptake by Roots and Stem
10-d old maize seedling (A) aluminum sample chamber; (B) neutron radiograph at ~70 µm pixel resolution illustrating roots distribution (0.2-1.6 mm); C) 3D tomographic reconstruction; (D) Timing of water uptake by plant components highlighted in (B) illustrating impact of solar radiation on rate of water flux in stem and ~0.5 mm first and second order roots.
This study provides direct evidence for root-mediated hydraulic redistribution of soil water to rehydrate drier roots
Notes: In soil plant systems, water variations were recorded in 10-day old Zea. mays (corn) seedlings following irrigation with D2O. D2O displacement of existing H2O within the plant vascular system, or influx of D2O into previously dry tissue or soil resulted in changes in image intensity through time. The dynamic internal demand for water between soil, individual roots, stems and leaves was thus readily visualized and quantified illustrating the value of this technique to assess root-shoot responses to changing environmental conditions.

## Slide 45: Ancient Craft Skills meet Modern Characterization
K. Ryzewski (PI), S. Herringer, H. Z. Bilheux, J.-C. Bilheux, B. Sheldon

## Slide 46: Thank you
Courtesy of
Krysta Ryzewski, Wayne University
Susan Herringer, Brian Sheldon, Brown University