## Slide 1
Achievement
Cryostat at HB3A Four-Circle Diffractometer
The cryostat has been tested at HB3A FCD and the data collected by the current 5x5cm Anger camera at 1.5 K were well fitted and consistent with the structure measured at 4 K with the regular CCR. The background is about 50% higher because the detector mask is off. The successful commissioning helps to estimate the background level for the large area detector upgrade and proved the capability of diffraction at ultra-low temperature, one of the benefits of the large area detector upgrade.
A successfully single crystal neutron diffraction was done at 1.5 K and magnetic order occurs at 3.4 K. Both magnetic and nuclear structures are well refined.
Katie Andrews, Cory Fletcher, Tyler White, Eric Stringfellow, Mike Hittman, Jon Smith, Gary Taufer
Huibo Cao, Yan Wu, and Bryan Chakoumakos
Structure refinement: red and black dots are observed and calculated squared structure factors respectively. The blue dots show difference. RF -factor : 2.258; Chi2(Intens): 1.330
Nuclear structure Magnetic structure Magnetic Ordering parameter
(½ ½ ½ )
Cory Fletcher
Tyler White
Katie Andrews