## Slide 1: X-ray Detection
Antonino Miceli (amiceli@aps.anl.gov)
August 7, 2016
NX School

## Slide 2: Outline
Counting versus Integrating
Indirect versus Direct detection
  Area detectors using scintillators
    Large area for diffraction (low spatial resolution, 50-100 mm)
    Small area for imaging (high spatial resolution, ~ 1 mm)
  Pixel array detectors (e.g., Pilatus, Pixirad)
Energy resolving detectors (i.e., spectroscopic detectors)
  Counting and direct detection
  Measuring the energy of photons
  Silicon diodes

## Slide 3: How do you detect x-rays?
Need to convert to something that you can measure
  Electrons… Q = CV
    Directly (x-rays  electrons)
      Ion chambers, pixel array detectors (e.g., Pilatus)
    Indirectly (x-rays  optical photons  electrons)
      Scintillators + Optics + photomultiplier/CCDs
  Temperature
    DT = Eg / (Heat Capacity)
      Superconducting calorimeters

## Slide 4: Counting versus Integrating Goal: Measure total flux of photons.
Analog signal from detector
Pulse Height ~ Energy
Integrating
Digitizer
Dark Current
Counting
threshold
?
?

## Slide 5: Counting versus Integrating
Counting
  Single photon counting
    Scintillator counting detectors (e.g., Cyberstar), avalanche photodiodes (APDs)
    Pilatus (counting pixel array detectors)
    Energy-resolving Detectors (Silicon or Germanium diode detectors)
  Deadtime limitations!!!
  Dark current rejected with a sufficiently high threshold.
  Easier to realize detectors
Integrating
  Signal accumulates
  CCDs, a-Si flat panel detectors
  No deadtime limitations
  Read noise and dark current are issues to consider
  Harder to realize detectors

## Slide 6: Counting detectors - Shaping time
Response time of detector
  In most detectors, the user can change this via software.
Gain is usually associated with longer shaping time.
Longer shaping time improves the energy resolution (to a point)
  But reduced the total count rate throughput.
  Energy resolution is important for certain techniques (e.g., XRF, XANES, XAS)
Analog response of a counting detector to the same x-ray photon energy

## Slide 7: Deadtime limitations for counting detectors
Analog pulses
Discriminator output (Digital)
Lower Threshold

## Slide 8: Deadtime
As you increase the input count rate (ICR), does the output count rate (OCR) follow linearly?
  The longer the shaping time, the lower the ICR before deviating from linearity.
When to worry (rule of thumb)?
  Rate > 1 / (2 x t)
  e.g., t ~ 1 ms  500 kHz
  Should probably calibrate to get an accurate flux measurement!
    Very important when you are looking for small changes (e.g., XAFS wiggles or XMCD)
Input count rate (ICR)
Output count rate (OCR)

## Slide 9: Deadtime for synchrotron (pulsed source)
Depends on the fill pattern and speed of the detector
Counting detectors can detect at most one photon per bunch
D.A. Walko, D.A. Arms, E.C. Landahl, J. Synchro. Rad. 15 (2008) 612; http://dx.doi.org/10.1107/S0909049508022358
NaI
YAP
APD
Fast detector (shorter shaping time)

## Slide 10: What fill pattern pattern will you be using?
Hybrid singlet and 324 bunch mode each 2 weeks a run.
Hybrid singlet useful for certain timing experiments.
  Not great for high count rate experiments (e.g., XANES, XAFS, XRF)

## Slide 11: Indirectly (x-rays  optical photons  electrons) Charge Coupled Devices (CCDs)
Optical detectors are everywhere in our lives… camera phones, etc.
CCDs are integrating detectors. No dead-time issues, but read noise and dark current

## Slide 12: Indirectly (x-rays  optical photons  electrons) Charge Coupled Devices (CCDs) + x-ray scintillators
Thickness of scintillator dictates the usable energy range

## Slide 13: Indirectly (x-rays  optical photons  electrons) Charge Coupled Devices (CCDs) + x-ray scintillators
With demagnification for large area detectors
  Diffraction ( < 30 keV)
x-rays
sample
Fiber Optic Taper (Optical photons)
(1 – 3 De-Magnification)
CCD
Scintillator (e.g., Gd2O2S)
Spatial resolution ~ 100 mm
No deadtime correction
  Integrating
Calibrations
  Dark Subtraction
  Spatial Distortion
  Spatial gain variations

## Slide 14: Indirectly (x-rays  optical photons  electrons) Amorphous Silicon Flat Panel + x-ray scintillators
Used at higher energies ( > 50 keV); integrating detectors.
Thin film transistor (TFT) technology (a-Si photo-sensors) allows large area detectors
  Cheaper than CCDs, but higher noise and less dynamic range!
  Typically faster than CCDs.
~ 120 cm
Sector 1

## Slide 15: Spatial/Geometric Distortions
Calibration can be done with a mechanical mask (“holey plate”) or a known x-ray sample
Input
Output

## Slide 16: Lagging/Ghosting/Memory effects
Can effect main types of detectors (e.g., GE & PerkinElmer flat panels, CdTe-pixel array detectors

## Slide 17: Direct Detection (x-rays  electrons)

## Slide 18: Directly (x-rays  electrons) Pixel Array Detectors (e.g., Pilatus)
Dectris Ltd

## Slide 19: Directly (x-rays  electrons) Pixel Array Detectors (e.g., Pilatus)
Pilatus is a digital PAD (photon counting)
CMOS readout chip (i.e., Application Specific Integrated circuit, ASIC)
PSI/SLS Detector Group
AGIPD pixel

## Slide 20: Directly (x-rays  electrons) Integrating Pixel Array Detectors
You can design the CMOS readout in anyway you like.
  e.g., with an integrating front end.
CSPAD at LCLS
Gruner et al.,

## Slide 21: Directly (x-rays  electrons) Pixel Array Detectors (e.g., Pilatus)
Each pixel is a single photon counting detectors!
Thus has count rate limitations
487 x 195 pixels (172 mm)
8.3 cm x 3.3 cm Area
Count Rate ~ 1 MHz/pixel
20-bit counter/pixel
5ms readout (Frame Rate = 200 Hz)
320 micron thick Silicon sensor
Gateable & electronic shutter
Lower Level Discriminator only
Brönnimann et al. @ PSI in Switzerland (Dectris)

## Slide 22: Commercial Photon Counting Detectors
Many photon counting detectors available today
  Dectris
  Medipix 3 detectors (X-Spectrum, Quantum Detectors, ASI)
  ImXPAD
  Pixirad
  Pixel sizes: 55-75 microns
  Frame rates up to 2-3 kHz.
Most vendors offer high-Z sensors (i.e., CdTe or GaAs).
  Pixirad seems to be the most reliable CdTe detector for a reasonable price.
  CdTe has “memory” problems
    Helped with cooling and HV switching
http://quantumdetectors.com/merlin/
https://www.dectris.com/
http://www.pixirad.com
http://www.x-spectrum.de/

## Slide 23: Photon counting, speed and dynamic range
Pilatus specs
  487 x 195 pixels (172 microns)
  Count Rate ~ 1 MHz/pixel
  20-bit counter/pixel
  Frame Rate = 200 Hz
  Gateable & electronic shutter
  Lower Level Discriminator only
You do not have 20-bit dynamic range @ 200 Hz!!!
Frame Rate (Hz) | Dynamic Range
1 Hz | 106 (20-bits)
10 Hz | 105
100 Hz | 104
1000 Hz | 103
DR decreases with speed!

## Slide 24: Directly (x-rays  electrons) Pixel Array Detectors – Photon Counting Threshold
Where to set the threshold?
Is there an “optimal” threshold?
Threshold

## Slide 25
Pixels
Photon Counting - Threshold –charge sharing

## Slide 26: Photon Counting - Threshold –charge sharing
If threshold is too high, then you under count events (effectively a small pixel)
If threshold is too low, then you double count events
“Optimal” threshold is 50% of beam energy
  Unless you need to reject fluorescent background.
B. Schmitt et al

## Slide 27: Energy resolving detectors (i.e., spectroscopic detectors)

## Slide 28: Energy Resolving Detectors – Counting (aka Energy Dispersive Detectors) (aka Spectroscopic Detectors) (aka XRF detector)
x-ray beam
Copper
XRF detector
Cu Ka
Cu Kb
Spectrum

## Slide 29: Fluorescence (XRF) Measures…
Abundance (ppm level) and spatial correlations of heavy elements
Also used for XAFS/XANES and XMCD
Elemental Compositions of Comet 81P/Wild 2 Samples Collected by Stardust (Flynn et al 2006) Solid-phases and desorption processes of arsenic within Bangladesh sediments (Polizzotto et al 2006)
Sea Coral
Matt Newville, 13-id
A link between copper and dental caries in human teeth identified by X-ray fluorescence elemental mapping (Harris et al 2008) Levels of Zinc, Selenium, Calcium, and Iron in Benign Breast Tissue and Risk of Subsequent Breast Cancer (Cui et al 2007)
XRF Image of a cell
Barry Lai, 2-id

## Slide 30: Spectroscopic Detectors
x-ray
X-Ray Energy ~ # of e-h pair
( 3.67 eV are need to produce 1 e-h pair for Silicon!!)
( Silicon or Germanium )

## Slide 31: In reality, we use Silicon Drift Diodes …
Noise scale like the area of the anode!
  Thus make the anode small

## Slide 32: Spectroscopic Detectors – Pulses to Histograms
3 ms
1 V
Pulse Height
Fe-55 Source on a scope.
MCA = histogram (e.g., 2048 channels)
Mn Kb
Mn Ka
6.5 keV
(Mn Kb)
5.9 keV
(Mn Ka)

## Slide 33: 4-element Silicon Drift Diode
4-element SDD (SII Nano Inc)
Energy resolution ~ 150eV
Peak-to-Background Important!!!
  Usually signal is buried here!
  Recombination, incomplete charge capture, etc.
Need to cool for -30C (via ThermoElectric Cooler)
Optimizing detector geometry for trace element mapping by X-ray fluorescence, Sun, Y et al Ultramicroscopy 152 (2015) 44–56

## Slide 34: Trade-off between count rate and energy resolution!!!
Shorter shaping time (length of pulse) means more count rate, but less energy resolution.
  Depends on your experiment.
Remember that the fill pattern will effect your deadtime correction, especially if you want to anything quantitative!

## Slide 35: Closing thoughts
Take a moment to think about the detector you are using.
Do you understand the detection mechanisms?
What are the limitations of the detectors?
  How does this effect the design of your experiment?
  Have the detector settings been optimized for your experiment?
What are the nonidealities of the detectors?
  Deadtime, spatial distortion, gain variations
  Can they be corrected?
You want to understand these things early in your experiment before you go home with “bad” data!

## Slide 36: Extra slides

## Slide 37: Dynamic Range
We typically have weak signals next to strong signals in the same image.
For counting detectors, dynamic range is set by how big your counter is (e.g., Pilatus has a 20-bit counter in each pixel)
For integrating detectors (e.g., CCD, a-Si flat panels), the dynamic range is set by how much charge each pixel can storage and the noise level.

Input Flux
Detector Signal
Level
DR = S/
S
Input Flux
Detector Signal
Level
DR = S
S
1
Gruner et al.,

## Slide 38: Spatial Resolution: Indirect vs. Direct
Indirect detection area detectors resolution is limited by the thickness of the scintillator because of optical blurring
Direct detection detectors have single-pixel resolution
Tungsten Knife Edge
GE a-Si Flat Panel (CsI)
Pilatus Silicon
X-ray
Optical Photons
Scintillator

## Slide 39: Directly (x-rays  electrons) Ion Chambers
Integrating detectors… ion current ~ x-ray flux
Used to monitor beam intensity
Used to normalize data to the beam intensity (“I0”)
Also used for transmission XAS measurements.

## Slide 40: Scintillator Thickness and Optimal Energy Range
Simple Energy-Dependence Model for Indirect Detection Area Detectors
Signal Size
marCCD/165
500 mm CsI
40 60 80 100 120 140
Energy (keV)
Number of X-ray Photons for S/N ratio = 1
8
7
6
5
4
3
2
GE a-Si Flat Panel
40 mm Gd2O2S
10
2
10 30 50 70 90
Energy (keV)

## Slide 41: Spectroscopic Detectors – Signal Chain
3 ms
1 V
Pulse Height
Digitizer
Pulse Height Analyzer
Multi-Channel Analyzer
Histogram
x-ray
diode
Preamp
Shaping Amp
Energy ~ P.H. ~ channel #