## Slide 1: Neutron Generation and Detection Neutron optics and Instrumentation
Lee Robertson
Instrument & Source Division
Oak Ridge National Laboratory
18th National School on Neutron and X-ray Scattering
July 30 – August 13, 2016
August 1, 11:00-12:00 Lecture

## Slide 2: What is a Neutron Scattering Instrument?
Neutron scattering experiments measure the number of neutrons scattered by a sample as a function of the wavevector change (Q) and the energy change (E) of the neutron.
What do we need to accomplish this?
  A source of neutrons
  A method for selecting the wavevector of the incident neutrons (ki)
  A very interesting sample
  A method for determining the wavevector of the scattered neutrons (kf)
  A neutron detector
Incident Wavevector
Source
Sample
Detector
Scattered Wavevector
ki
kf

## Slide 3: Why Not Just Build a Universal Neutron Scattering Instrument That Can Do Everything We Need?
Two types of sources (continuous and pulsed)
Two methods for determining the neutron wavevector, k (time-of-flight and diffraction)
Two types of scattered neutrons (elastic and inelastic)
Two types of interactions between the neutrons and the sample (nuclear and magnetic)
Wide range of length scales driven by the science
The energy of the neutron is coupled to its wavelength and velocity:
l2(Å2) ~ 81.81/E(meV) and v2(m2/s2) ~ 191313×E(meV)
S(Q,E) the scattering properties of the sample depend only on Q and E, not on the neutron wavelength(l)
Message: Many different types of neutron scattering instruments are needed because the accessible Q and E ranges depend on the neutron energy and because the resolution and detector coverage have to be tailored to the science for such a signal-limited technique.
Incident Wavevector
Source
Sample
Detector
Scattered Wavevector
ki
kf

## Slide 4: Pulsed vs Continuous Apples vs Oranges
Spallation Neutron Source (pulsed)
High Flux Isotope Reactor (continuous)
The peak neutron production of the SNS is about 10x that of the HFIR
The HFIR neutron production is about 15x the time averaged production of the SNS
Neutrons Produced (n/s)
Time (ms)

## Slide 5: Neutron Scattering Instruments at Continuous Sources Are Typically Based on Diffraction Techniques
Source
Sample
Detector
“White” Beam
Monochromatic Beam
Transmitted Beam
Scattered Beam
Scattering
Angle
Monochromator
Crystal
Transmitted Beam
Bragg’s Law
Momentum Transfer
Transmitted
Beam
Detector
Analyzer
Crystal
ki
kf

## Slide 6: Neutron Scattering Instruments at Pulsed Sources Are Typically Based on Neutron Time-of-Flight Techniques
Sample
Source
Detectors
Fermi
Chopper
Time and Distance
v(1.8Å) = 2187m/s
TOF(s) = D(m)/v(m/s) = [D(m)×λ(Å)]/3956.0339
D=20m, TOF(1Å)=0.005s, TOF(2Å)=0.010s, DTOF=0.005s
l(Å) = (3956.0339 * TOF(s)) / D(m)

## Slide 7: Neutron Optics
The following neutron optical components are typically used to construct a neutron scattering instrument
  Monochromators / Analyzers: Monochromate or analyze the energy of a neutron beam using Bragg’s law
  Choppers: Define a short pulse of neutrons or select a small band of neutron energies
  Guides / Mirrors: Allow neutrons to travel large distances without suffering intensity loss
  Polarizers / Spin Manipulators: Filter and manipulate the neutron spin
  Collimators: Define the direction of travel of the neutrons
  Detectors: Neutron position (and arrival time for TOF) is recorded. Neutrons are typically detected via secondary ionization effects.

## Slide 8: Liouville's Theorem
In the geometrical-optics the propagation of neutrons can be represented as trajectories in a six-dimensional phase space (p, q), where the components of q are the generalized coordinates and the components of p are the conjugate momenta.
Simply stated, Liouville's Theorem says that phase space volume is conserved.
Translation: It costs flux to increase resolution and it costs resolution to increase flux.
There is no way to win!
Parallel
Converging

## Slide 9: Instrument Resolution
Uncertainty in the neutron wavelength and direction limit the precision that Q and E can be determined
For scattering, the uncertainty comes from how well ki and kf can be determined
For TOF, the uncertainty primarily comes from not knowing the exact start time for each neutron
The total signal observed in a scattering experiment is proportional to the phase space volume within the elliptical resolution volume – the better the resolution, the lower the count rate
Figure borrowed from Roger Pynn

## Slide 10: Choppers and Velocity Selectors
TOF Timing Diagram
Distance
1
2
3
4
5
6
7
Frame
Disk Chopper
Fermi Chopper
Velocity Selector
Astrium
J-PARC

## Slide 11: Neutron Guides
80m Guide for HRPD at J-PARC
Fabricated by Swiss Neutronics
m=0.1 deg/Å
Multichannel Curved Guide
Fabricated by Swiss Neutronics
Guide Installation at ISIS

## Slide 12: Polarizers and Spin Manipulators
Spherical Neutron Polarimetry
POLI-HEiDi at FRMII
3He Cell
Unpolarized
Neutron Beam
Polarized
Neutron Beam
3He Spin
Filters
Heussler Monochromator
AlCuMn
Polarizing Supermirrors
Larmor Precession
Flipper

## Slide 13: Neutron Optics: Focusing
Doubly focusing Cu monochromator at the ILL
Double focusing “Popovici” monochromator. The vertical curvature is fixed while the horizontal curvature is variable by bending stacks of thin silicon wafers. The gain is achieved both by spatial focusing and ‘wavelength focusing’.
Virtual Source
Small
Aperture
Mono
Sample

## Slide 14: Neutron Optics: Focusing
Focusing Mirrors :
  Develop a nested advanced KB mirror system to make a more compact assembly and to achieve the highest performance.
  Identify applications where focusing optics can replace neutron guides and offer better performance.
90 m

## Slide 15: Elastic Neutron Scattering Instruments
Small Proteins
Micelles Colloids Vesicles
Critical Phenomena
Membranes
Liquids and Glasses
Crystal,
Fiber,
and
Magnetic
Structures
Polymers and
Metals Alloys Zeolites
Magnetic Films
and Domains
Biological Structure
Viruses
SANS
REFLECTOMETERS
DIFFRACTOMETERS
Elastic instruments include:
  Powder diffraction
  Single Crystal diffraction
  SANS (typical)
  Reflectometry
  Used to determine the average structure of materials (i.e. how the atoms are arranged)
Based on diagram by Roger Pynn
Longer Length Scales

## Slide 16: TOF Powder Diffractometer: POWGEN (SNS)
Sr2Fe1.5Mo0.5O6, Electrode Material for Solid Oxide Fuel Cells
J. Am. Chem. Soc., 134, p6826 (2012)
For POWGEN D = 64.5m

## Slide 17: Small Angle Neutron Scattering (SANS)
TmNi2B2C Vortex Lattice
PHYS REV B 86, 144501 (2012)

## Slide 18: Magnetism Reflectometer (SNS)
PHY REV B 84, 245310 (2011)

## Slide 19: Neutron Imaging
Carbon foam matrix in a Li battery
(H. Bilheux and S. Voisin)
X-ray image
Neutron Image

## Slide 20: Inelastic Neutron Scattering Instruments
Inelastic instruments include:
0.1
1
10
100
d(Å)
Used to study dynamics such as phonons, magnons, and diffusion (i.e. what the atoms are doing)
Neutron Induced
Excitations
Hydrogen
Modes
Molecular
Vibrations
Intenerate
Magnets
Crystal
Fields
Spin
Waves
Lattice
Vibrations
Coherent
Modes
Electron-Phonon
Interactions
Molecular
Motions
Aggregate Motion
Polymers and Biological Molecules
Critical Scattering
Diffusion
Molecular
Reorientation
Based on diagram by Rex Hjelm
Slower
Motions
Longer Length Scales
Direct Geometry TOF Spectrometers
Indirect Geometry TOF Spectrometers
Triple-Axis Spectrometers
Backscattering Spectrometers
Neutron Spin-Echo Spectrometers

## Slide 21: SEQUOIA: A Direct Geometry TOF Spectrometer at the SNS
Quantum oscillations of nitrogen atoms in uranium nitride
Nature Communications v3, p1124 (2012)

## Slide 22: Triple-Axis Spectrometer
Lattice Dynamics of PbTe
PHYS REV B 86, 085313 (2012)

## Slide 23: BASIS: An Inverted Geometry Backscattering Spectrometer
Study of water diffusion on single-supported bilayer lipid membranes by QENS
Bai M., EPL 98, 48006 (2012)

## Slide 26: Advanced Neutron Optics: IMAGINE
CG-4 Guide
CTAX
Mono
Entrance Slit 1
Flat Mirror
Beamstop/Slit
Flight Tube
Multimirrors
Entrance Slit 2
INSTRUMENT ENCLOSURE
Diffractometer
Drop-In Apertures
Beamstop/Slit
Monitor
Elliptical
Mirror 1
Elliptical
Mirror 2
  Short wavelength filter @ 2.0, 2.78, and 3.33 Å
  Long wavelength filter @ 3.0, 4.0, and 4.5 Å
  Elliptical mirrors for focusing the beam
4 mm
2 mm
beam focus at sample position

## Slide 27: Future Instruments: SESANS/SERGIS
Real space correlation lengths up to 20 microns (and beyond?)
Does not require tight collimation for high resolution
Can be used to probe the in-plane correlations of thin films and interfaces.
Spin-Echo Scattering Angle Measurement:
The neutron spin precesses through two parallelogram-shaped magnetic fields in opposite directions. For scattered neutrons the path-length through the two parallelograms is different resulting in a net change in the spin angle.
+B
sample
-B
analyzer
detector
polarizer
off-specular
specular
horizon
l
q
sample
Using SERGIS the specular
and off-specular scattering
can be distinguished

## Slide 28: Concluding Remarks
Instrument design is driven by the needs of the scientific community coupled with the source capabilities along with advances in neutron optics and detectors.
In the near term instrument development will be primarily focused on:
  Focusing optics
  Polarization
  Detectors
  Instrument development infrastructure (computer simulations)
  New techniques and applications