Freezer-dryer- Auto mode instructions

[For specialty applications of manual mode- consult the instrument user manual]

Start-up Procedure

- Place samples in -80 oC freezer in desired vessel. The sample container should be 2 – 3 times larger than the samples. Seal top of vessel with perforated parafilm to allow solvent to be removed from sample during freeze-drying

- Meanwhile, get the freeze dryer ready for operation.

- Check that the collector chamber is clean and dry. Replace lid

- Make sure that the sample valves are positioned to the vent position (Bevel points upwards)

- Power up the freeze dryer, the power switch is on the left-hand side of the cabinet

- To run in auto mode, press the panel switch labeled REFRIDGERATION AUTO (see fig 1). This will initiate refrigeration and once -40oC is reached the vacuum will start.

Operating Procedure

- When the system vacuum level is at <0.133 mBar, samples can be added

- Connect the pre-frozen sample to a sample valve on the manifold using an adapter

- Turn the plastic knob to the vacuum position. The bevel points towards the sample port (facing down)

- Before adding a second sample, allow the vacuum level to return to <0.133 mBar

- To remove a sample, turn the plastic knob on the sample valve to the vent position (Bevel points up). This will close the valve and vent the sample container

Shut-down Procedure

- Release system vacuum by turning the plastic knob to the vacuum position (this is the open position with bevel facing down)

- Press the vacuum switch (see Fig 1) on the control panel to turn the vacuum off

- Press the refrigeration button (see fig 1) next to the illuminated LED to turn off the refrigeration system

- Turn off the Main power switch on the left-hand side of the cabinet

- Clean the collector chamber

- Pull the collector chamber drain hose out from the cabinet and remove the drain plug. Place hose in suitable container to catch liquids from the chamber

- Add a small amount of warm water to melt ice around coils. Once it has melted enough, the ice can be manually removed. [Important Note: The water should not cover more than the bottom 2 coils otherwise water may drain into vacuum pump]

- Flush collector chamber with water (see note above) and wipe dry

Figure 1. Front display of Freeze-dryer