1100 C Ashing Oven

1100 °C Ashing Oven

| Potential Hazards while performing these activities |  |
| --- | --- |
| Heat Hazard | This equipment gets extremely hot and can cause burns. |
| Inhalation Hazard | If liquids are heated, dangerous fumes may be released. This product contains refractory ceramic fiber, which can result in irritation of the skin, eyes, and respiratory tract. |
| Chemical Hazard | Chemicals may be used with this unit, spills may occur. |

| Preventions to reduce exposure to hazards: |  |
| --- | --- |
| Eye Protection | Wear Safety Glasses while operating. |
| Protective Gloves | Wear appropriate gloves for heat and chemical use. |
| Read Manual/Guide | Become familiar with this guide before operating. Read the manual if you need to use multiple set points. |

Note: This QSG is for single set point operations

| Setup Steps | Pictures |
| --- | --- |
| - Turn the furnace on by flipping the power switch. Press and HOLD for 3 seconds the “SET/ENT” button to display “modE rES”. |  |
| - Press and release the ‘SET/ENT’ button to display “PrG 0”. Press the ‘UP/RESET’ button to show the lower display value of “1”. |  |
| - Press and release the ‘SET/ENT’ button to select this new value and advance to the “SSP 25” display. |  |
| - Press and release the ‘SET/ENT’ button until the High Temperature Alarm Setpoint value is displayed as “A1”. Select an alarm setpoint 10 degrees C above the target setpoint to be selected. Press and release the ‘SET/ENT’ button to place this new value in the controller memory. |  |
| - “Hy1” will be displayed, this completes setup. Press and HOLD for 3 seconds the ‘SET/ENT’ button to exit this menu. |  |
| Operation | Pictures |
| - Press and hold for 3 seconds ‘SET/ENT’ button to display “modE rES”. Press and release the ‘UP’ button twice to select the display “LCL”. |  |
| - Press and release the ‘SET/ENT’ button once to select Local Mode. This selection causes the red indicator to illuminate beside “L” on the control panel. Note: you will hear a “click” when you complete this step. |  |
| - Use the ‘UP’ and ‘DOWN’ buttons to select the desired operating temperature setpoint. Press and release the ‘SET/ENT’ button once to register the setpoint value. |  |
| - Display will show measured temperature in the upper display, the present temperature setpoint in the lower display. |  |
| - This display and the buttons will remain active as long as power continues to the control module. Power interruptions will cause the controller to enter reset or standby mode. |  |
| - When finished, turn the power off by flipping the power switch and allow the furnace to cool to room temperature before opening. Use hot mitts when removing items from furnace. |  |