Mettler pH Meter

Mettler pH Meter

| Potential Hazards while performing these activities |  |
| --- | --- |
| Electrical Hazard | This equipment contains electrical units that may come into contact with skin or liquids. |
| Chemical Hazard | This equipment is used near potentially hazardous chemicals. |
| Corrosive Hazard | This equipment is used for measuring acids and bases. |
| Chocking Hazard | Because liquids are agitated, dangerous fumes may be released. Use of a fume hood may be required. |
| Slip Hazard | Liquids from the unit may fall on the floor, causing slippery conditions. |

| Preventions to reduce exposure to hazards: |  |
| --- | --- |
| Eye Protection | Wear safety glasses while operating. |
| Protective Gloves | Wear appropriate gloves for chemical usage. |
| Non-Slip, Closed Shoes | Wear non-slip, closed shoes to avoid spills. |
| Read Manual | Become familiar with this guide before operating. |

Note: - Make sure the calibration buffer range includes the pH of the samples. - The slope is automatically calculated after each buffer. Make sure the slope is >95%.

| Setup Steps | Pictures |
| --- | --- |
| - Setup the base and ensure the arm is centered. |  |
| - Ensure the electrode is plugged into the socket. |  |
| - Ensure power supply is plugged in. |  |
| - Turn the unit on. |  |
| Operation | Pictures |
| - Decide if manual or automatic endpoints will be used. Hold READ to switch from manual endpoints to automatic endpoints and vice versa. The A =Automatic endpoint. The radical without the A =Manual endpoint. |  |
| - Remove cap from electrode. |  |
| - Rinse electrode with de-ionized water. |  |
| - Place electrode in pH 7 buffer and stir. |  |
| - Press Start Calibration. Wait until the screen stops blinking if using the automatic endpoint. Press READ again if using the manual endpoint. |  |
| - Rinse with de-ionized water and wipe with a kim wipe. |  |
| - Place electrode in the next pH buffer for calibration. |  |
| - Repeat steps 5 and 6 for a 3-point calibration. |  |
| - Make sure the slope reads 95-105%. |  |
| - Rinse with de-ionized water and wipe with a kim wipe. Place electrode in sample. The pH meter will automatically start reading the pH for the sample after the 3-point calibration so work quickly. Press Read for manual mode. |  |
| - When finished, remove and rinse electrode with de-ionized water. Place electrode back into the filling solution and recap. |  |