## Slide 1: Polarization / CCR Cart
Polarization CCR tipping analysis

## Slide 2: Polarization CCR tipping analysis
The cart will be analyzed for two conditions.
  CCR extended
  CCR retracted
  Adding counterweight 20x17x4.5 inches, 433lbs
The cart wheels will be oriented to worst case scenario
Analysis done per UL 1678
Conclusion:
CCR extended and retracted analysis
  The cart is unstable to transport with CCR.
  Cart did not pass side loading criteria or 10° incline.
  Dynamic was acceptable
Added counterweight
  To stabilize the cart a 433 lb counterweight would need to be added to bottom of cart.
  With the counterweight all criteria is met per UL 1678
  The cart would need to be reanalyzed for the added weight
  Lift analysis would need to be reviewed

## Slide 3: Cart with CCR extended out
Center Of
Gravity
32.732
3.872
12.128
8.883
13.367
6.75°
15.18°
20.33°
22.21°
Fore-Aft
Side to Side

## Slide 4: Dynamic tip stability
OK
OK

## Slide 5: Tipping due to side loading
UL 1678 16.5.3
Will only analyze worst case, where the CG is closest to the side.
32.732
3.872
12.128
Cart will tip from applied force
40 lbs
422 lbs
48.00

## Slide 6: Tipping due to incline of 10°
422 lbs
Center of gravity
Incline plane 10°
The center of gravity is beyond point “M” and would tip over.

## Slide 7: Cart with CCR retracted in
32.732
11.582
4.418
Center of
gravity
Fore-Aft
Side to Side
8.883
13.367
15.18°
22.21°
7.69°
19.49°

## Slide 8: Dynamic tip stability
OK
OK

## Slide 9: Tipping due to side loading
48.00
11.582
4.418
UL 1678 16.5.3
Will only analyze worst case, where the CG is closest to the side.
Cart will tip from applied force
40 lbs

## Slide 10: Tipping due to incline of 10°
422 lbs
Incline plane 10°
The center of gravity is beyond point “M” and would tip over.

## Slide 11: Cart with CCR extended out and counterweight
5.102
10.898
21.41
9.164
13.086
13.40°
26.98°
31.43°
counterweight
23.17°
Fore-Aft
Side to Side

## Slide 12: Dynamic tip stability
OK
OK

## Slide 13: Tipping due to side loading
5.102
10.898
UL 1678 16.5.3
Will only analyze worst case, where the CG is closest to the side.
Cart will not tip from applied force
48.00
40 lbs

## Slide 14: Tipping due to incline of 10°
5.102
10.898
855 lbs
The center of gravity is behind point “M” and would not tip over.